UNIVERSITY OF BRISTOL SCHOOL OF CHEMISTRY RISK ASSESSMENT This form must be completed by a competent assessor for any procedure/system of work before an attempt is made at carrying out the procedure/system of work. Please refer to the instructions for making a Risk Assessment in the School of Chemistry Safety Manual (http://www.chm.bris.ac.uk/safety/nfrass.htm).

Name and Status of the Assessor: Dr. James Smith (Research Officer) Date: 10/9/18

Activity/procedure being assessed: Realignment and unrestricted optical access of Laser Raman System (S111)

Known or expected hazards associated with the activity: - Laser Radiation (Direct and scattered light) - Electric shock (High Voltages) - UV emission from laser - Organic solvents (highly flammable) - Silicon/diamond substrates

The risk of injury and its severity likely to arise from these hazards: - Alignment of the Raman system requires full unrestricted access to the optical system with the removal of elements of the safety infrastructure (i.e. defeated interlocks). All three lasers used within the Raman system are rated Class 3B and above, exposure of direct and scattered light is harmful to eyes and potentially harmful to skin. - The risk of injury from electric shock in normal operation is extremely low since hazardous high voltages are present only inside the laser power supplies and the Raman enclosure. Equipment casing is sealed and should never be opened. The risk is very low but potential severity is very high. - During operation the He-Cd UV laser produces quantities of emission that may be hazardous to human skin and eyes. Care should be taken to prevent exposure of skin and eyes to the laser emission. Wear laser glasses when appropriate. - The risk of injury due to a solvent fire is low to medium: the risk of ignition is minimal. The severity of this injury is likely to be low to medium in view of the primary application of solvents for cleaning samples which requires only very small quantities. The cleaning solvents usually employed are ethanol and acetone, so they pose only negligible toxic hazard. The potential severity of injury by methanol (sometimes used for cleaning optics) is greater due to its toxicity, but even smaller quantities are needed in this case, so that the overall impact is limited. - Substrates are normally handled with tweezers to avoid contaminating or damaging them; thus, the risk of cuts is low. Since the sharp objects in question are very small, the potential severity of a cut from them is inherently limited.

Who is at risk? Research workers using the apparatus. Only trained users should be in the lab while realignment is being carried out.

Measure to be taken to reduce the level of risk: - Only trained operators have access to the complete optical system. Care must be taken to eliminate all unwanted scattered and back-reflected light. No laser radiation should be emitted from the enclosed area. While aligning the lab should be clear of untrained personnel and laser safety curtains pulled shut. Appropriate laser glasses should be worn during alignment. All laser guards must be checked and maintained. - All electrical items must be operated with no hazardous voltages accessible to users, and repair and maintenance of equipment must be performed by a competent person, and with electrical power completely disconnected wherever technically possible. Metal casings and other exposed parts of electrical apparatus must be bonded to earth ground to guard against possible dangerous faults. Warning signs are placed on panels that prevent user access to high voltage parts, these panels should not be removed while the equipment is powered. - As per point 1 above. Additional precautions should be taken to avoid exposure of the laser to skin. - Cleaning of apparatus and optics must be performed using a minimal quantity of solvent in a well-ventilated area, away from potential sources of ignition. Spills must be cleaned up promptly to avoid later accidental ignition. Bulk solvents must be stored in a suitable solvent cabinet, and wash bottles are to be of the vapour-venting type and stored appropriately. Protective gloves should preferably be worn to avoid skin irritation, and this is mandatory when the solvent used is methanol due to its established percutaneous toxicity. - No further measures deemed necessary. If a cut is sustained, it will be sufficient to wash with water and apply a sticking plaster.

Training prerequisites: Untrained persons must not work on the experiment unsupervised.

Level of risk remaining: Slight, barring unforeseeable accidents.

Action to be taken in an emergency: IN CASE OF INJURY, perform First Aid if trained or contact a designated First Aider. (A current list is provided at the Porter’s Lodge). If necessary, also contact security on extension 112233 (0117 331 1223) to call for an ambulance. IN CASE OF FIRE, use a carbon dioxide fire extinguisher, two are located in the corridor near to the entrance to S111. If the fire is too large to deal with in this way, evacuate the affected area, raise the fire alarm, and contact security on extension 112233 (0117 331 1223) to call the fire brigade.

References: The following constitute an integral part of this risk assessment and must be reviewed along with the above. Standard School of Chemistry risk assessments (see http://www.chm.bris.ac.uk/safety/compform.htm): - Use of standard electrical equipment - Transport, Storage and Use of Solvents and other Flammable Liquids - Manual Handling - Visual Display Equipment

Signature of Assessor: Supervisor’s signature: J.A. Smith P.W. May