UNIVERSITY OF BRISTOL SCHOOL OF CHEMISTRY RISK ASSESSMENT This form must be completed by a competent assessor for any procedure/system of work before an attempt is made at carrying out the procedure/system of work. Please refer to the instructions for making a Risk Assessment in the School of Chemistry Safety Manual (http://www.chm.bris.ac.uk/safety/nfrass.htm).

Name and Status of the Assessor: James Smith (Research Officer) Date: 7/10/16

Activity/procedure being assessed: Secondary electron emission Test rig

Known or expected hazards associated with the activity: - Electric shock (High Voltages) and safe practice of using an electron beam gun - High Vacuum apparatus (implosion hazard) - Organic solvents (highly flammable) - Sample Handling

The risk of injury and its severity likely to arise from these hazards: - The risk of injury from electric shock in normal operation is extremely low since hazardous high voltages are only present inside the vacuum chamber and the electron beam power supply enclosure. Equipment casing is sealed and should never be opened. The electron gun must only be operated when the system is at high vacuum. The vacuum chamber and all electrical supplies must be earthed. The risk of electric shock is very low but potentially severe. - The risk of injury from implosion is very low as almost the entire apparatus is solidly constructed from stainless steel. The load lock door seal must be checked regularly and the level of vacuum constantly monitored. - The risk of injury due to a solvent fire is low to medium: the risk of ignition is minimal. The severity of this injury is likely to be low to medium in view of the primary application of solvents for cleaning samples and chamber fittings which requires only very small quantities. The cleaning solvents usually employed are ethanol and acetone, so they pose only negligible toxic hazard. The potential severity of injury by methanol (sometimes used for cleaning optics) is greater due to its toxicity, but even smaller quantities are needed in this case, so that the overall impact is limited. - Test samples are normally handled with tweezers to avoid contaminating or damaging them; thus, the risk of cuts is low. Since the sharp objects in question are very small, the potential severity of a cut from them is inherently limited.

Who is at risk? Research workers using the apparatus. There is a very remote chance that other people nearby in the laboratory may be affected in the unlikely event of a major incident.

Measure to be taken to reduce the level of risk: - All electrical items must be operated with no hazardous voltages accessible to users, and repair and maintenance of equipment must be performed by a competent person, and with electrical power completely disconnected wherever technically possible. Metal casings and other exposed parts of electrical apparatus must be bonded to earth ground to guard against possible dangerous faults. - Regular checks of the vacuum system are essential. - Cleaning of apparatus must be performed using a minimal quantity of solvent in a well-ventilated area, away from potential sources of ignition. Spills must be cleaned up promptly in order to avoid later accidental ignition. Bulk solvents must be stored in a suitable solvent cabinet, and wash bottles are to be of the vapour-venting type and stored appropriately. Protective gloves should preferably be worn to avoid skin irritation, and this is mandatory when the solvent used is methanol due to its established percutaneous toxicity. - No further measures deemed necessary. If a cut is sustained, it will be sufficient to wash with water and apply a sticking plaster.

Training prerequisites: As required in the references. Training in the particular safety aspects of operating the apparatus is intrinsic in the training required in order to operate it in other respects, i.e. without damaging equipment and in order to obtain meaningful results. Untrained persons must not work on the experiment unsupervised.

Level of risk remaining: Minimal, barring unforeseeable accidents.

Action to be taken in an emergency: IN CASE OF INJURY, perform First Aid if trained or contact a designated First Aider. (A current list is provided at the Porter’s Lodge). If necessary, also contact security on extension 112233 (0117 331 1223) to call for an ambulance. IN CASE OF FIRE, use a carbon dioxide fire extinguisher, two are located in the corridor near to the entrance to S111. If the fire is too large to deal with in this way, evacuate the affected area, raise the fire alarm, and contact security on extension 112233 (0117 331 1223) to call the fire brigade.

References: The following constitute an integral part of this risk assessment and must be reviewed along with the above. Standard School of Chemistry risk assessments (see http://www.chm.bris.ac.uk/safety/compform.htm): - Use of standard electrical equipment - Use of Reduced Pressure or Vacuum - Transport, Storage and Use of Solvents and other Flammable Liquids - Manual Handling - Visual Display Equipment

Signature of Assessor: Supervisor’s signature: J.A. Smith N.A. Fox